## Slide 1: Quality Assurance Measures and External Standards

## Slide 2: Subject Test Sequence
  Enter subject, officer, and arrest information
  Diagnostic Check
  Blank Test
  Internal Standard Check
  External Dry Gas Standard Check
  Blank Test
  Subject Sample
  Blank Test
  External Dry Gas Standard Check
  Blank Test
  All components must complete successfully to obtain a breath alcohol result.

## Slide 3: QA Components of Every Subject Test
Diagnostic Check
Blank Test
Internal Standard Check
External Dry Gas Standard Check

## Slide 4: Diagnostic Check
Voltages and temperatures of the instrument are checked to ensure they fall within the approved ranges.
Result will read “PASSED”.
Subject test will stop if this does not pass.

## Slide 5
Notes: Change screen shot to stop at diagnostic check

## Slide 6: Blank Test
A measurement is taken after the AMBIENT ZEROING phase of the operation. This ensures the sample chamber is clear before the subject sample can enter.
Must read “0.000” or the subject test will stop automatically.
Occurs another 3 times during the subject test sequence.

## Slide 7
Notes: Change screen shot to stop at diagnostic check

## Slide 8: Internal Standard Check
1 mm thick piece of quartz moved into the light path. A measurement of the quartz plate is stored at the time of Adjustment.
Quartz plate helps ensure the accuracy of the DataMaster DMT has not changed since the last adjustment.
Internal standard is measured with every test and compared to the stored value.
“Verified” on the test strip indicates that the check passed.

## Slide 9
Notes: Change screen shot to stop at diagnostic check

## Slide 10: External Standard Check
External dry gas standard: nitrogen and ethanol
A sample of the dry gas standard is pushed through the gas line into the sample chamber, is analyzed, and the results are displayed.
The external standard delivers a known quantity of ethanol to the DataMaster both before AND after the subject sample.
The purpose is to ensure the instrument is accurately recognizing and measuring ethanol.

## Slide 11
External Dry Gas Standards
Commercially-prepared dry gas standards
Provided to each agency by the Crime Lab

## Slide 13: External Dry Gas Standards
Automatic internal regulator (no need to turn On or Off)
Every tank will have concentration of 0.080 g/210 L

## Slide 14: Barometric Pressure Adjustment
Cylinders are a closed, pressurized system.
Must adjust the ethanol concentration for pressure at time of test (expansion of compressed gas).
DataMaster has an internal barometer and measures pressure at beginning of the subject test sequence to compensate (automatic).
Subject samples are not adjusted because humans are not a closed, pressurize system.

## Slide 15: Barometric Pressure Adjustment
At Standard pressure (29.92) external standards = 0.080
Lower pressure (<29.92)
External standards
< 0.080
Higher pressure (>29.92)
External standards
> 0.080

## Slide 16: Tolerance of the External Dry Gas Standard
Each external dry gas reading must be +/- 0.005 or less of the target value adjusted for barometric pressure.

## Slide 17
The two external standard readings must be +/- 0.005 of the “target at” value.

## Slide 18
The SUBJECT SAMPLE is NOT adjusted for pressure.
ONLY the external dry gas standard is adjusted for pressure.

## Slide 19: Completed Test Strip
Flow and alcohol are graphed
Subject Sample lists a number
No status messages

## Slide 20: QA of the Breath Program
Verifications of Calibration (VOC)
Adjustment
Calibration

## Slide 21: Verification of Calibration (VOC)
A quality assurance measure performed at least every 60 days and each time a dry gas external standard is changed.
A Verification of Calibration consists of two main parts:
Diagnostic test
5 tests of the external dry gas standard

## Slide 22: Verifications of Calibration
VOC DOES NOT calibrate the DataMaster (only done by the Crime Lab)
60 days or less requirement is from…
13 AAC 63.100 (c)
At intervals not to exceed 60 days, the accuracy of the calibration of a breath test instrument must be verified.
Notes: 13 AAC 63.100 (c) At intervals not to exceed 60 days, the accuracy of the calibration of a breath test instrument must be verified. The verification of calibration must be performed by the scientific director or by a qualified person designated by the scientific director. A written report of the verification of calibration shall be made by the person who performed that verification. That report shall be submitted to the scientific director, if the scientific director did not prepare it. Upon receipt of the written report or after preparing the written report, and if the report verifies that the calibration of the breath test instrument is accurate, the scientific director shall issue a certification for the breath test instrument. The certification is valid for 60 days after the day the verification of calibration was performed. The scientific director may require verification of calibration of a breath test instrument at more frequent intervals.

## Slide 24: Verification of Calibration
The Verification will be performed by the instrument automatically and uploaded by the Crime Lab at least every 60 days.
It will be performed on these dates at noon. It will not interrupt a subject test.
  Jan 3rd
  Feb 21st
  April 11th
  May 30th
  July 18th
  Sept 5th
  October 24th
  December 12th

## Slide 25: Verification of Calibration
A VOC will also be performed automatically after an external standard cylinder is changed.
Breath test supervisors and operators may initiate a VOC on the DataMaster DMT using the DMT Icon and pressing “VOC Test”.
Use this option for:
  Changing instrument location in your building
  After maintenance or repair

## Slide 26: Quality Assurance of the VOC
Standard Deviation requirement on VOC (measure of variability of readings)
Accepted range for the external standard (+/- 0.005) from the adjusted target value (accuracy)

## Slide 28: Accuracy
This is the 0.005 tolerance of the readings of the external standard from the adjusted target value.
Applies to external standard measurements from the subject test sequence AND
Applies to Verifications of Calibration (VOCs)

## Slide 29: Precision
Measured by the Standard Deviation
A low standard deviation indicates that the data points tend to be very close together (good), whereas high standard deviation indicates that the data points are spread out over a large range (bad).
Accepted Standard Deviation is 0.0030 or less
Applies only to the VOC and internal lab testing, not to the subject test.

## Slide 30: Examples that will fail the S.D.
0.085
0.080
0.080
0.075
0.076
Std Dev = 0.0063
0.085
0.080
0.080
0.079
0.079
Std Dev = 0.0042

## Slide 31: Example that will pass the S.D.
0.081
0.080
0.080
0.079
0.079
Std Dev = 0.0014

## Slide 33: VOC Process
Two components of the VOC complete successfully.
File is saved to memory of the DataMaster.
VOC file is uploaded by the crime lab via the modem line or by the Breath Test Supervisor.

## Slide 34: VOC Review Process
VOC is printed out and reviewed by one of the scientists – this review includes:
  Ensure diagnostic test passed
  Lot number and expiration date of gas cylinder are correct
  All five external standard readings are +/- 0.005
  Standard deviation is 0.0030 or less
This review is documented with their initials and date on the bottom right corner of VOC.

## Slide 36: VOC Process
VOC is reviewed again by the Scientific Director.
VOC is notarized.
Signature certifies the instrument for continued evidentiary use in the State from the date of the VOC.
Signed VOCs are posted on the Crime Lab website and are sortable by instrument serial number.

## Slide 38: Unsuccessful VOCs
If any portion of the VOC does not pass, the instrument will not allow any other testing to be performed.
The DataMaster will only allow another VOC to be performed.

## Slide 39: Calibration
Instrument Calibration is the process that the Breath Alcohol discipline uses to demonstrate that the individual breath instrument is working properly.
The DataMaster DMT Calibration process includes voltage checks, linearity testing, functional tests for triggering status messages, minimum volume determination, and a non-drinking subject test.

## Slide 40: Calibration
Occurs before a DataMaster leaves the laboratory
  After arrival at the lab when returned from manufacturer
  After maintenance has been performed by lab
  Preventative Maintenance (At least every 4 years for all instruments)

## Slide 41: Adjustment
DataMaster DMT adjustment is a single point adjustment (0.100 g/210 L).
The Crime Lab performs adjustment of the DataMaster DMT using a wet bath ethanol simulator solution and a blank simulator solution. The blank simulator solution is 500 mL of deionized water.
Simulator: Sealed, temperature-controlled jar with hoses attached.

## Slide 42: Review
Adjustment of breath test instruments is performed by Breath Test Supervisors.
(True or False)
False: Adjustment is only performed by the instrument manufacturer or the crime lab staff and must be followed by instrument calibration.

## Slide 43: Review
External standards will be out of range at +/- 0.004 of the target value adjusted for barometric pressure.
(True/False)
False: Standards are out of range above
+/- 0.005.

## Slide 44: Review
At what PSI will the instrument…
Warn you to change your tank?
Not let you run a test?
(2 answers)
1. 150 psi – Warning to change your tank
2. 100 psi – Won’t let you run a test

## Slide 45: Review
What two tests are part of a VOC? (2 answers)
5 tests of the external standard
Diagnostic Test